PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

# Online Phosphate Analyser with predictive diagnostics

A high-performance analyser designed for precise measurement of ortho-phosphate in water using the molybdovanadate yellow colorimetric method. The instrument features automatic zero-point calibration without the need for standard solutions, ensuring consistent accuracy by compensating for sample background color.

Key capabilities include:

- Grab sample capability for direct correlation with laboratory reference measurements

- Self-cleaning functionality for reduced maintenance

- Rapid measurement cycles – as frequent as every 5 minutes

- Built-in self-diagnostic routines for enhanced reliability

- Humidity sensor to detect internal leaks and trigger safe shutdown procedures

- Radial piston pump for robust fluid handling

- Universal digital controller requirement for full operational integration

Advanced predictive diagnostics extend operational uptime by monitoring internal components and alerting users to upcoming maintenance needs. When paired with a compatible controller, users can view real-time status indicators for each sensor with predictive capabilities.

Additionally, the analyser supports remote monitoring via any browser-enabled device, offering visibility into sensor

health, diagnostics, and maintenance schedules—helping reduce risk and minimise downtime.

# Technical Data

| Online Phosphate Analyser for low range |  |
| --- | --- |
| Measuring range | 0.015 - 5.0 mg/L PO4-P |
| Detection limit | 0.015 mg/L PO4-P |
| Accuracy | With Standard Solution: 2% of the measured value + 0.015 mg/L |
| Repeatability | With Standard Solution: 0.7% of the measured value + 0.005 mg/L |
| Reagent consumption | 500 mL/month |
| Online Phosphate Analyser for mid range |  |
| Measuring range | 0.05 - 15 mg/L PO4-P |
| Detection limit | 0.05 mg/L PO4-P |
| Accuracy | With Standard Solution: 2% of the measured value + 0.05 mg/L |
| Repeatability | With Standard Solution: 2% of the measured value + 0.05 mg/L |
| Reagent consumption | 1000 mL/month |
| Online Phosphate Analyser for high range |  |
| Measuring range | 1 - 75 mg/L PO4-P |
| Detection limit | 1 mg/L PO4-P |
| Accuracy | With Standard Solution: 2% of the measured value + 1.0 mg/L |
| Repeatability | With Standard Solution: 2% of the measured value + 1.0 mg/L |
| Reagent consumption | 1000 mL/month |

| Online Phosphate Analyser (in general, for all ranges) |  |
| --- | --- |
| Measurement method | Photometric (Ortho-phosphate ions react with vanadate-molybdate reagent and form a yellow dye.) |
| Measuring interval | 10, 15, 20 or 30 minutes (user adjustable) |
| pH range | 5 - 9 pH |
| Permissible Chloride range | ≤ 1000 mg/L Cl- |
| Sample quality | Ultra filtrated or comparable |
| Flow rate | 0.5 - 20.0 L/h |
| Sample pressure | 2.5 bar (0.25 mPa) maximum |
| Sample temperature | 4 - 40 °C |
| Special remarks | The liquid level in the basin must be below the bottom of the analyser. |
| Operating temperature range | -20 - 45 °C; 95% relative humidity, non-condensing |
| Storage conditions | -20 - 60 °C; 95% relative humidity, non-condensing |
| Number of channels | 1 or 2 |
| Enclosure waterproof rating | Rating: IP55, NEMA 3R UL50E Material: PUR 66 |
| Controller compatibility | Digital SC Controller. Controllers are sold separately. |
| Power supply | Mains power supplied by SC Controller or power box. Analyser and heated drain tubing: 115 VAC or 230 VAC |
| Cable length | Standard: 2 m cable fixed from the edge of the enclosure Extension cables are available as an option (data cable up to 50 m, power cable up to 5 m). |
| Dimensions (H x W x D) | 991 mm x 575 mm x 425 mm |
| Weight | Approximately 45 kg without chemicals |

Scope of delivery

[ ] (delete what is not needed)

# Analytical Instrument

[ ] Online Phosphate Analyser for use with external filtration system

[ ] Online Phosphate Analyser with sample supply pump for use with integrated filtration system

[ ] Online Phosphate Analyser with sample supply pump for use with integrated filtration system, rough samples

# Measuring Ranges

[ ] Low Range (0.015 - 5.0 mg/L PO4-P) [ ] Mid Range (0.05 - 15 mg/L PO4-P)

[ ] High Range (1 - 75 mg/L PO4-P)

# Number of Channels:

[ ] 1 channel [ ] 2 channels

# Instrument Options

[ ] Flow detection

[ ] Automatic grab sample and sample verification

[ ] Flow detection + automatic grab sample and sample verification

# Accessories

[ ] Keys for analyser enclosure (1 replacement pair) [ ] Heated drain/connecting hose, 2 m, 230 V

[ ] Mounting kit for analyser incl. fastening and angle bracket, screws [ ] Accessories for continuous sample (1-/2-channel)

[ ] Power extension cable for digital controller, 5 m, 115-230 VAC

[ ] Door for analyser indoor housing including. 4 stickers

[ ] Door for analyser enclosure including 4 instrument labels [ ] Power supply box for analyser, with EU plug

[ ] Power connection box for analysers

# Mounting Hardware

[ ] Extension pipe, 1.0 m, made of stainless steel [ ] Extension pipe, 1.8 m, made of stainless steel [ ] Rail mounting for analyser with controller

[ ] Stand mounting kit, suitable for 1 analyser + 1 digital controller [ ] Stand mounting kit, suitable for 1 analyser

[ ] Rail mounting for analyser without digital controller

[ ] Weather guard for digital controller for outdoor installation

# Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

[ ] External sample filtration system for Phosphate Analyser

NOTE: Sample filtration is strongly recommended in all waste water process areas. If you are specifying a two stream analyser in this design you should specify two filtration units, one for each sample stream!

Brand:	Hach

Product:	NP6000sc Online Phosphate Analyser with Prognosys predictive diagnostics