## Release Notes
Package # | Application Code Version | Release Date (use text for month)
4 | 1 | December 2025
List changes: initial release

## Languages
Section 4: Firmware Information
Languages Implemented:
| English (0): | X
| German (1): | X
| Spanish (2): | X
| French (3): | X
| Italian (4): | X
| Netherlands (5): | X
| Danish (6): | X
Swedish (7):
| Chinese (8): | X
| Polish (9): | X
Japanese (10):
Korean (11):
| Portuguese (12): | X
Slovak (13):
Russian (14):
Hungarian (15):
Bulgarian (16):
Rumanian (17):
Czech (18):
Turkish (19):
Finnish (20):
Greek (21):
Serbian (23):
Slovenian (24):
Thai (25)

## Data Log Information
Section 6: Data Log Information
Number of Data Log Channels (16 Maximum): List the log channel information in the following format: (add more rows as needed)
Channel | Register | Tag Name
0 | 40001 | measurement_channel_1
1 | 40002 | measurement_channel_2
2 | 40011 | tgflow_channel_1
3 | 40013 | tgflow_channel_2
4 | not used | not used
5 | not used | not used
6 | not used | not used
7 | not used | not used
8 | not used | not used
9 | not used | not used
10 | not used | not used
11 | not used | not used
12 | not used | not used
13 | not used | not used
14 | not used | not used
15 | not used | not used
16 | not used | not used

## Event Log Information
Section 7: Event Log Information
Number of Event Log Channels (256 Maximum): List the log channel information in the following format: (add more rows as needed)
Number | Number of Tags | Event Text Name | Description
0 | 2 | ERROR | An error has bee detected.
49
50
1 | 2 | WARNING | Warning!
117
50
2 | 2 | PROCESS EVENT | event logger process event
118
50
3 | 2 | REMINDER | event logger reminder
119
50
4 | 0 | POWER ON EVENT | Power on event
5 | 0 | SENSOR POWER | Sensor power failure
6 | 0 | SET DATE/TIME | Set date and time.
7 | 2 | CODE VERSION | Embedded software version
| 22 | Application
| 22 | Application
8 | 2 | DRIVER VERSION | Device Driver Version
| 21 | Driver Version
| 21 | Driver Version
9 | 2 | SCRIPT VERSION | Script version
| 20 | Script Version
| 20 | Script Version
10 | 1 | SOFTWARE VERS | Software version
| 22 | Application
11 | 3 | FIELD CALIB. FULL | Full calibration
| 135 | Calibration Result
293
293
12 | 4 | FIELD CALIB. FULL | Full calibration
| 135 | Calibration Result
293
293
293
13 | 3 | FIELD CALIB. FULL | Full calibration
300
301
302
14 | 6 | FIELD CALIB. FULL | Full calibration
300
301
302
303
304
305
15 | 3 | FIELD CALIB. ZERO | Field zero calibration
| 135 | Calibration Result
293
302
16 | 4 | FIELD CALIB. ZERO | Field zero calibration
| 135 | Calibration Result
293
302
305
17 | 3 | FIELD CALIB. ZERO | Field zero calibration
| 135 | Calibration Result
278
293
18 | 3 | AUTO CAL | Automatic calibration
| 135 | Calibration Result
293
302
19 | 2 | CLEANING | Cleaning
146
159
20 | 3 | GRAB SAMPLE | Grab sample
| 2 | NH4N Ch1
| 2 | NH4N Ch1
| 2 | NH4N Ch1
21 | 3 | VALIDATION | Validation
| 2 | NH4N Ch1
| 2 | NH4N Ch1
| 2 | NH4N Ch1
22 | 3 | AIR TIGHTNESS TEST | Air tightness test
113
113
113
23 | 2 | NODE CONNECTION STATUS | Connection Status
111
103
24 | 2 | HEATING PERIOD CHANGED | Heating Period Changed
57
58
25 | 1 | MEAS. CLIMAT. FAILED! | Measurement chamber climatization failure
| 120 | Chamber temperature
26 | 2 | DRAIN IS BLOCKED! | Drain is blocked
113
113
27 | 1 | DEVICE IS TOO COLD! | Device is too cold
| 116 | Housing temperature
28 | 1 | DEVICE TOO HOT! | Device is too hot
| 116 | Housing temperature
29 | 1 | DEVICE THAWING! | Device Thawing
| 116 | Housing temperature
30 | 1 | DEVICE NOT HEATING! | Device not heating
| 116 | Housing temperature
31 | 2 | FLOW CHANNEL 1 | Flow channel 1
| 6 | FLOW CHANNEL 1
152
32 | 2 | FLOW CHANNEL 2 | Flow channel 2
| 7 | FLOW CHANNEL 2
153
33 | 0 | STANDARD 1 | Calibration value Standard 1
34 | 0 | STANDARD 2 | Calibration value standard 2
35 | 0 | CLEANING SOLUTION | Cleaning solution
36 | 2 | WORKFLOW | Workflow
148
149
37 | 2 | DEVICE MODE CHANGE | Device Mode Change
150
151
38 | 3 | COUNTER RESET | Service - Counter reset
158
156
157
39 | 1 | SET DATE/TIME | Set date and time.
142
40 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 10 | LOCATION
| 10 | LOCATION
41 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 51 | Location Ch1
| 51 | Location Ch1
42 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 52 | Location Ch2
| 52 | Location Ch2
43 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 39 | Serial Number
| 39 | Serial Number
44 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 17 | MEAS UNITS
| 17 | MEAS UNITS
45 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
134
134
46 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 43 | Factor Ch1
| 43 | Factor Ch1
47 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 45 | Offset Ch1
| 45 | Offset Ch1
48 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
18
18
49 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 286 | Measurement Interval
| 286 | Measurement Interval
50 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 62 | Calibration start time
| 62 | Calibration start time
51 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 59 | Calibration Interval
| 59 | Calibration Interval
52 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 61 | Cleaning start time
| 61 | Cleaning start time
53 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 60 | Cleaning Interval
| 60 | Cleaning Interval
54 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
248
248
55 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
102
102
56 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
57
57
57 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
58
58
58 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
47
47
59 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
48
48
60 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
66
66
61 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
138
138
62 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
152
152
63 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
63
63
64 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 53 | Number of measurements Ch1
| 53 | Number of measurements Ch1
65 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 54 | Number of measurements Ch2
| 54 | Number of measurements Ch2
66 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 55 | Discard measurements Ch1
| 55 | Discard measurements Ch1
67 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
| 56 | Discard measurements Ch2
| 56 | Discard measurements Ch2
68 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
64
64
69 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
65
65
70 | 3 | CH1 NO SAMPLE! | No Sample Channel 1
244
113
230
71 | 3 | CH2 NO SAMPLE! | No Sample for Channel 2
244
113
230
72 | 7 | CONFIGURATION | Configuration
| 19 | Measurement Interval
| 59 | Calibration Interval
| 62 | Calibration start time
| 60 | Cleaning Interval
| 61 | Cleaning start time
134
263
73 | 4 | DAILY TEMPERATURES | Daily Temperatures
| 116 | Housing temperature
| 116 | Housing temperature
| 116 | Housing temperature
| 116 | Housing temperature
74 | 2 | DAILY PRESSURE LEVELS | Daily Pressure Levels
113
113
75 | 2 | DAILY SAMPLING STATUS | Daily Sampling Status
| 79 | Filter module status
137
76 | 3 | BOTTLE LEVELS | Bottle Levels
143
146
145
77 | 5 | BOTTLE LEVELS | Bottle Levels
143
143
144
146
145
78 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
296
296
79 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
295
295
80 | 3 | PARAMETER CHANGED | Parameter was changed by user
161
264
264

## Warning List Information
Section 9: Warning List Information
Number of Warnings (32 Maximum): List the Warning List information in the following format:
Register | Bit # | Description
49983 | 0 | 6-MONTH MAINT. OVERDUE
49983 | 1 | AIRFILTER PADS OVERDUE
49983 | 2 | SOILED WINDOWS
49983 | 3 | PISTON HEAD OVERDUE
49983 | 4 | WARNING
49983 | 5 | CAL ZERO POOR
49983 | 6 | FILTRMOD CLEAN OVERDUE
49983 | 7 | FILTER MOD. OVERDUE
49983 | 8 | SAMPLE TUB. OVERDUE
49983 | 9 | SAMPLE HEAD OVERDUE
49983 | 10 | COMPRESSOR OVERDUE
49983 | 11 | MEAS OUT OF RANGE
49983 | 12 | CLEANING 1&2 LEVEL LOW
49983 | 13 | STAND. BLANK LEVEL LOW
49983 | 14 | CH1 SAMPLE FLOW LOW
49983 | 15 | CH2 SAMPLE FLOW LOW
49984 | 0 | AIR TIGHTNESS FAILED
49984 | 1 | REAGENT LEVEL LOW
49984 | 2 | CAL DATA INCORRECT
49984 | 3 | CLOSE DEVICE DOOR!
49984 | 4 | CALIBRATION IS OFF
49984 | 5 | CURRENT DIAGNOSTIC
49984 | 6 | DRIP PAN MISSING!
49984 | 7 | LED LEVEL LOW
49984 | 8 | REAGENT&ACID LEVEL LOW
49984 | 9 | not used
49984 | 10 | not used
49984 | 11 | not used
49984 | 12 | not used
49984 | 13 | not used
49984 | 14 | not used
49984 | 15 | not used

## Error List Information
Section 8: Error List Information
Number of Errors (32 Maximum): List the Error List information in the following format:
Register | Bit # | Description
49985 | 0 | ELECTRONIC ERROR!
49985 | 1 | DRAIN IS BLOCKED!
49985 | 2 | MEAS. CLIMAT. FAILED!
49985 | 3 | CAL ZERO INCORRECT
49985 | 4 | DEVICE LEAKAGE!
49985 | 5 | DEVICE NOT HEATING!
49985 | 6 | DEVICE TOO HOT!
49985 | 7 | DEVICE THAWING!
49985 | 8 | PRESSUR.SENS. DEFECT!
49985 | 9 | TRACE HEAT. FAILED!
49985 | 10 | SOILED WINDOWS
49985 | 11 | DEVICE IS TOO COLD!
49985 | 12 | TEMP.SENS. DEFECTIVE!
49985 | 13 | LED FAULTY
49985 | 14 | CH1 NO SAMPLE!
49985 | 15 | CH2 NO SAMPLE!
49986 | 0 | AIR TIGHTNESS FAILED
49986 | 1 | STRAYLIGHT TOO HIGH
49986 | 2 | CAL DATA INCORRECT
49986 | 3 | not used
49986 | 4 | not used
49986 | 5 | not used
49986 | 6 | not used
49986 | 7 | not used
49986 | 8 | not used
49986 | 9 | not used
49986 | 10 | not used
49986 | 11 | not used
49986 | 12 | not used
49986 | 13 | not used
49986 | 14 | not used
49986 | 15 | not used

## Process Event List Information
Section 10: Process Events List Information
Number of Process Events (30 Maximum): List the process event information in the following format:
Register | Bit # | Description
49987 | 0 | WARMING UP...
49987 | 1 | UPDATING SENSOR...
49987 | 2 | COOLING DOWN...
49987 | 3 | THERM. READJUSTMENT...
49987 | 4 | PREPARING NEXT PROCESS
49987 | 5 | PREPARING OPER. MODE
49987 | 6 | PREPARING MAINT. MODE
49987 | 7 | MAINTENANCE MODE
49987 | 8 | CALIBRATING...
49987 | 9 | CLEANING...
49987 | 10 | CLEANING ETL...
49987 | 11 | GRAB SAMPLE...
49987 | 12 | VALIDATION...
49987 | 13 | MEASURING...
49987 | 14 | MEASURING CH1...
49987 | 15 | MEASURING CH2...
49988 | 0 | DISCARD MEASUREMENT
49988 | 1 | DISCARD MEAS. CH1
49988 | 2 | DISCARD MEAS. CH2
49988 | 3 | PREPARING SAMPLE...
49988 | 4 | PREPUMPING...
49988 | 5 | PREPUMPING SAMPLE...
49988 | 6 | PREPUMPING REAGENT...
49988 | 7 | PREPUMPING STANDARDS..
49988 | 8 | INITIALIZING FLUIDIC..
49988 | 9 | FLUIDIC SELF-TEST...
49988 | 10 | not used
49988 | 11 | not used
49988 | 12 | not used
49988 | 13 | not used
49988 | 14 | Prohibited, internal use only
49988 | 15 | Prohibited, internal use only

## Classified Error
Section 14: Classified Error
List Classified Errors.
Register | Error # | Error | Description
49930 | 0 | Measurement Calibration Error | Faulty calibration detected
49930 | 1 | Electronic Adjustment Error | Faulty electronic calibration/settings
49930 | 2 | Cleaning Error | Error in cleaning cylce detected
49930 | 3 | Measuring Module Error | Error in measuring module detected
49930 | 4 | System Re-initialization Error | Inconsistent settings detected, reset to factory settings
49930 | 5 | Hardware Error | Faulty hardware detected
49930 | 6 | Internal Communication Error | Internal communication error detected
49930 | 7 | Humidity Error | Excessive hunidty detected
49930 | 8 | Temperature Error | Excessive temperature detected
49930 | 9 | Reserved for Future Use
49930 | 10 | Sample Warning | Error in sample feed detected
49930 | 11 | Questionable Calibration Warning | Accuracy of previous calibration inadequate
49930 | 12 | Questionable Measurement Warning | Accuracy of previous measurement inadequate/out of range
49930 | 13 | Safety Warning | Safety equipment error detected
49930 | 14 | Reagent Warning | Reagent warning, e.g fill level < min detected
49930 | 15 | Maintenance Required Warning | Service request detected

## Classified Status
Section 14: Classified Status (Register 1/2)
List Classified Status.
Register | Bit # | Status | Description
49931 | 0 | Calibration in progress | Calibration in progress, measuremet value not up to date
49931 | 1 | Cleaning in progress | Cleaning in progress, measurement value not up to date
49931 | 2 | Service/Maintenance menu | Device in Service mode
49931 | 3 | Common Error | Error in device
49931 | 4 | Measurement 0 Quality Bad | Measurement accuracy is not within specified limits
49931 | 5 | Measurement 0 Low Limit | Measurement value falls short of the specified range
49931 | 6 | Measurement 0 High Limit | Measurement value exceeds the specified range
49931 | 7 | Measurement 1 Quality Bad | Measurement accuracy is not within specified limits
49931 | 8 | Measurement 1 Low Limit | Measurement value falls short of the specified range
49931 | 9 | Measurement 1 High Limit | Measurement value exceeds the specified range
49931 | 10 | Measurement 2 Quality Bad | Measurement accuracy is not within specified limits
49931 | 11 | Measurement 2 Low Limit | Measurement value falls short of the specified range
49931 | 12 | Measurement 2 High Limit | Measurement value exceeds the specified range
49931 | 13 | Measurement 3 Quality Bad | Measurement accuracy is not within specified limits
49931 | 14 | Measurement 3 Low Limit | Measurement value falls short of the specified range
49931 | 15 | Measurement 3 High Limit | Measurement value exceeds the specified range
49932 | 0 | Reserved for future use | not supported
49932 | 1 | Reserved for future use | not supported
49932 | 2 | Reserved for future use | not supported
49932 | 3 | Reserved for future use | not supported
49932 | 4 | Measurement 4 Quality Bad | not supported
49932 | 5 | Measurement 4 Low Limit | not supported
49932 | 6 | Measurement 4 High Limit | not supported
49932 | 7 | Measurement 5 Quality Bad | not supported
49932 | 8 | Measurement 5 Low Limit | not supported
49932 | 9 | Measurement 5 High Limit | not supported
49932 | 10 | Measurement 6 Quality Bad | not supported
49932 | 11 | Measurement 6 Low Limit | not supported
49932 | 12 | Measurement 6 High Limit | not supported
49932 | 13 | Measurement 7 Quality Bad | not supported
49932 | 14 | Measurement 7 Low Limit | not supported
49932 | 15 | Measurement 7 High Limit | not supported
Section 14: Classified Status (Register 3/4)
List Classified Status.
Register | Bit # | Status | Description
49933 | 0 | Reserved for future use | not supported
49933 | 1 | Reserved for future use | not supported
49933 | 2 | Reserved for future use | not supported
49933 | 3 | Reserved for future use | not supported
49933 | 4 | Measurement 8 Quality Bad | not supported
49933 | 5 | Measurement 8 Low Limit | not supported
49933 | 6 | Measurement 8 High Limit | not supported
49933 | 7 | Measurement 9 Quality Bad | not supported
49933 | 8 | Measurement 9 Low Limit | not supported
49933 | 9 | Measurement 9 High Limit | not supported
49933 | 10 | Measurement 10 Quality Bad | not supported
49933 | 11 | Measurement 10 Low Limit | not supported
49933 | 12 | Measurement 10 High Limit | not supported
49933 | 13 | Measurement 11 Quality Bad | not supported
49933 | 14 | Measurement 11 Low Limit | not supported
49933 | 15 | Measurement 11 High Limit | not supported
49934 | 0 | Reserved for future use | not supported
49934 | 1 | Reserved for future use | not supported
49934 | 2 | Reserved for future use | not supported
49934 | 3 | Reserved for future use | not supported
49934 | 4 | Measurement 12 Quality Bad | not supported
49934 | 5 | Measurement 12 Low Limit | not supported
49934 | 6 | Measurement 12 High Limit | not supported
49934 | 7 | Measurement 13 Quality Bad | not supported
49934 | 8 | Measurement 13 Low Limit | not supported
49934 | 9 | Measurement 13 High Limit | not supported
49934 | 10 | Measurement 14 Quality Bad | not supported
49934 | 11 | Measurement 14 Low Limit | not supported
49934 | 12 | Measurement 14 High Limit | not supported
49934 | 13 | Measurement 15 Quality Bad | not supported
49934 | 14 | Measurement 15 Low Limit | not supported
49934 | 15 | Measurement 15 High Limit | not supported

## Modbus Register
Tag Name | Register # | Data Type | Length | R/W | Units(U) | Parameter(P) | Discrete Range | Min/Max Range | Description
NH4N Ch1 | 40003 | Float | 2 | R | MEAS UNITS | PO4-P |  | 1500
NH4 Ch1 | 40005 | Float | 2 | R | MEAS UNITS | PO4 |  | 1500
NH4N Ch2 | 40007 | Float | 2 | R | MEAS UNITS | PO4-P |  | 1500
NH4 Ch2 | 40009 | Float | 2 | R | MEAS UNITS | PO4 |  | 1500
FLOW CHANNEL 1 | 40011 | Float | 2 | R | mL/h | FLOW |  | 20000
FLOW CHANNEL 2 | 40013 | Float | 2 | R | mL/h | FLOW |  | 20000
Remote Control | 40025 | Unsigned Integer | 1 | R/W |  |  |  | 65535 | Remote control
Ammonium Ch1 | 40028 | Float | 2 | R | MEAS UNITS |  |  | 1500
Electrode Voltage | 40030 | Float | 2 | R | mV |  |  | 5000
Script Version | 40039 | Unsigned Integer | 1 | R |  |  |  | 9999
Driver Version | 40040 | String | 6 | R
Application | 40046 | Float | 2 | R |  |  |  | 3.402823E+38
Version Valve Unit | 40050 | String | 6 | R
Version Sample Unit | 40056 | String | 6 | R
Version Piston Pump Unit | 40062 | String | 6 | R
Version Senosr Unit | 40068 | String | 6 | R
Version Housing Unit | 40074 | String | 6 | R
Version AC Unit | 40080 | String | 6 | R
Version Dosing Unit | 40086 | String | 6 | R
HW index Control Unit | 40092 | Unsigned Integer | 1 | R |  |  |  | 65535
HW index Valve Unit | 40093 | Unsigned Integer | 1 | R |  |  |  | 65535
HW index Sample Unit | 40094 | Unsigned Integer | 1 | R |  |  |  | 65535
HW index Piston Pump Unit | 40095 | Unsigned Integer | 1 | R |  |  |  | 65535
HW index Sensor Unit | 40096 | Unsigned Integer | 1 | R |  |  |  | 65535
HW index Housing Unit | 40097 | Unsigned Integer | 1 | R |  |  |  | 65535
HW index AC Unit | 40098 | Unsigned Integer | 1 | R |  |  |  | 65535
HW index Dosing Unit | 40099 | Unsigned Integer | 1 | R |  |  |  | 65535
Serial Number | 40100 | String | 6 | R
Factor Ch1 | 40117 | Float | 2 | R/W |  |  |  | 2
Factor Ch2 | 40119 | Float | 2 | R/W |  |  |  | 2
Offset Ch1 | 40121 | Float | 2 | R/W | MEAS UNITS |  |  | 50
Offset Ch2 | 40123 | Float | 2 | R/W | MEAS UNITS |  |  | 50
Location Ch1 | 40129 | String | 8 | R/W
Location Ch2 | 40137 | String | 8 | R/W
Number of measurements Ch1 | 40145 | Unsigned Integer | 1 | R/W |  |  |  | 10
Number of measurements Ch2 | 40146 | Unsigned Integer | 1 | R/W |  |  |  | 10
Discard measurements Ch1 | 40147 | Unsigned Integer | 1 | R/W |  |  |  | 10
Discard measurements Ch2 | 40148 | Unsigned Integer | 1 | R/W |  |  |  | 10
Calibration Interval | 40151 | Unsigned Integer | 1 | R/W |  |  | 0/1/3/4
Cleaning Interval | 40152 | Unsigned Integer | 1 | R/W |  |  | 0/1/2/3/4
Cleaning start time | 40153 | Unsigned Integer | 1 | R/W |  |  | 0/1/2/3/4/5/6/7/8/9/10/11/12/13/14/15/16/17/18/19/20/21/22/23
Calibration start time | 40154 | Unsigned Integer | 1 | R/W |  |  | 0/1/2/3/4/5/6/7/8/9/10/11/12/13/14/15/16/17/18/19/20/21/22/23
6-Month maintenance | 40160 | Integer | 1 | R | days |  |  | 32767
Range reagent solution | 40161 | Integer | 1 | R | days |  |  | 32767
Range cleaning solution | 40162 | Integer | 1 | R | days |  |  | 32767
Range standard solution | 40163 | Integer | 1 | R | days |  |  | 32767
Range blank solution | 40163 | Integer | 1 | R | days |  |  | 32767
Range air filter pads | 40164 | Integer | 1 | R | days |  |  | 32767
Range membrane and electrolyte | 40165 | Integer | 1 | R | days |  |  | 32767
Range electrode | 40166 | Integer | 1 | R | days |  |  | 32767
Range piston pump head | 40167 | Integer | 1 | R | days |  |  | 32767
Range reagent pump | 40168 | Integer | 1 | R | days |  |  | 32767
Range acid pump | 40169 | Integer | 1 | R | days |  |  | 32767
Filter module status | 40170 | Unsigned Integer | 1 | R | % |  |  | 100
Range Cleaning filter module | 40171 | Integer | 1 | R | days |  |  | 32767
Range new filtration module | 40172 | Integer | 1 | R | days |  |  | 32767
Range pump tubings | 40173 | Integer | 1 | R | days |  |  | 32767
Range pump head | 40174 | Integer | 1 | R | days |  |  | 32767
Range compressor | 40175 | Integer | 1 | R | days |  |  | 32767
Operation time | 40176 | Unsigned Integer | 1 | R | days |  |  | 65535
Housing temperature | 40209 | Float | 2 | R | C |  |  | 125
Chamber temperature | 40212 | Float | 2 | R | C |  |  | 125
Measurement chamber heater power | 40214 | Float | 2 | R | % |  |  | 100
Measured standard 1 | 40223 | Float | 2 | R | mV |  |  | 5000
Measured standard 2 | 40225 | Float | 2 | R | mV |  |  | 5000
Calibration Result | 40242 | Unsigned Integer | 1 | R |  |  | 0/1/2
Last calibration date | 40243 | Time2 | 2 | R
Slope | 40263 | Float | 2 | R |  |  |  | 5000
Offset | 40265 | Float | 2 | R |  |  |  | 5000
| 40383 | Float | 2 | R |  |  |  | 3.402823E+38
| 40386 | Unsigned Integer | 1 | R/W |  |  | 0/1/2/3/4
LED 1 operation time | 40408 | Float | 2 | R | h |  |  | 0
LED 2 operation time | 40410 | Float | 2 | R | h |  |  | 0
| 40451 | Float | 2 | R | mV |  |  | 0
| 40453 | Float | 2 | R | mA |  |  | 0
| 40466 | Float | 2 | R/W |  |  |  | 1.05
| 40468 | Float | 2 | R/W |  |  |  | 0.1
| 40473 | Float | 2 | R |  |  |  | 3.402823E+38
| 40475 | Float | 2 | R |  |  |  | 3.402823E+38
| 40477 | Float | 2 | R |  |  |  | 3.402823E+38
| 40479 | Float | 2 | R |  |  |  | 3.402823E+38
| 40481 | Float | 2 | R |  |  |  | 3.402823E+38
| 40483 | Float | 2 | R |  |  |  | 3.402823E+38